## Dashboard
Medical Equipment Maintenance Tracker
Total Tickets |  | New Tickets | In Progress | Completed
8 |  | 3 | 2 | 2
Status Summary
Status | Count
New | 3
In Progress | 2
Awaiting Parts | 1
Completed | 2
Priority Breakdown
Priority | Count
Critical | 2
High | 3
Medium | 2
Low | 1

## Maintenance Log
Medical Equipment Maintenance & Fault Tracker
Ticket ID | Request Date | Equipment Name | Serial Number (SN) | Department / Room | Issue Description | Reported By | Priority | Status | Assigned Engineer | Completion Date | Cost (USD) | Action Taken / Notes
REQ-2026-001 | 2026-09-01 | Ventilator | VENT-9921 | ICU | Pressure sensor error | Dr. Ahmed Khalil | Critical | Completed | Eng. Muhanad Ali | 2026-09-02 | 150 | Replaced sensor and recalibrated system
REQ-2026-002 | 2026-09-03 | Mobile X-Ray Machine | XRAY-4410 | Radiology | Main power switch unresponsive | Tech. Jasim | High | In Progress | Eng. Omar Farooq |  | 0 | Inspecting power supply board
REQ-2026-003 | 2026-09-05 | ECG Machine | ECG-1029 | Emergency Room | Signal cable damage | ER Duty Nurse | Medium | Awaiting Parts | Eng. Sarah Sadiq |  | 80 | New cable ordered from vendor
REQ-2026-004 | 2026-09-07 | Defibrillator | DEF-5532 | Operating Room | Battery not holding charge | Dr. Samar Fadhil | Critical | New | Eng. Muhanad Ali |  | 0 | Scheduled for initial inspection
REQ-2026-005 | 2026-09-08 | Infusion Pump | INF-8812 | VIP Ward | Touchscreen unresponsive | Head Nurse | Low | Completed | Eng. Khalid Abdullah | 2026-09-09 | 45 | Recalibrated touch screen interface
REQ-2026-006 | 2026-09-10 | MRI Scanner | MRI-0019 | MRI Suite | Cooling system over-temperature | Dr. Yasser Al-Ali | Critical | In Progress | External Service Team |  | 1200 | Vendor engineer onsite investigating chiller
REQ-2026-007 | 2026-09-11 | Dialysis Machine | DIA-7720 | Hemodialysis Unit | Minor leak in water inlet pipe | Dept Maintenance Tech | High | Completed | Eng. Sarah Sadiq | 2026-09-11 | 20 | Replaced inlet tube and verified pressure
REQ-2026-008 | 2026-09-12 | Infant Incubator | INC-3301 | NICU | Temperature fluctuation detected | Dr. Amal Nouri | High | New | Eng. Omar Farooq |  | 0 | Scheduled for testing
REQ-2026-009 | 2026-09-13 | ESG | 123456 | PICU | paper jam | Dr. Ahmed Khalil | Medium | New | Eng. Noor |  | 0 | Scheduled for testing
[10 empty] | Total Cost: | 1495

## Equipment Inventory
Medical Equipment Inventory Database
Equipment Code | Equipment Name | Serial Number (SN) | Manufacturer / Brand | Department / Location | Purchase Date | Warranty Status
EQ-01 | Ventilator | VENT-9921 | Hamilton Medical | ICU | 2023-05-15 | Active
EQ-02 | Mobile X-Ray Machine | XRAY-4410 | Siemens | Radiology | 2021-11-10 | Expired
EQ-03 | ECG Machine | ECG-1029 | GE Healthcare | Emergency Room | 2024-01-20 | Active
EQ-04 | Defibrillator | DEF-5532 | Mindray | Operating Room | 2022-08-05 | Active
EQ-05 | Infusion Pump | INF-8812 | B. Braun | VIP Ward | 2023-02-14 | Active
EQ-06 | MRI Scanner | MRI-0019 | Philips | MRI Suite | 2019-04-30 | AMC Contract
EQ-07 | Dialysis Machine | DIA-7720 | Fresenius | Hemodialysis Unit | 2022-12-01 | Active
EQ-08 | Infant Incubator | INC-3301 | Dräger | NICU | 2020-09-18 | Expired
EQ-09 | Infant Incubator | INC-3302 | Dräger | NICU | 2020-09-19 | Expired
EQ-10 | ESG | 123456 | Mindray | PICU | 2025-09-15 | Active